**Installation Notes**

MX2 Firmware 1.0.11


Prior to installing, it is recommended to review the MX2 Firmware 1.0.11 Release Notes. The following are the instruction to install the firmware update.




# Accessing MX2 Firmware


The website is the primary means for delivering firmware. If a CD is desired, one can be requested by contacting DSI Customer Service (CService@datasci.com).

1. Navigate to [www.datasci.com](http://www.datasci.com)
2. Browse to **Resources | DSI Gateway Portal**.
3. Login using your **User Name** and **Password**. If you have forgotten or do not have a User Name and Password, the ability to request a new password or register at the site is available.
4. Select **Software Updates**.
5. Click the **MX2 1.0.11.zip** to download the contents to a desired location.
6. Extract the files by **right-clicking** the .zip file and choosing **Extract**.


# Installing MX2 Firmware

1. Launch **Ponemah**.
2. Select the **Hardware** **|** **MX2 Diagnostics…***Note*: If using *Ponemah v5.20*, select **Hardware | Diagnostics…**

1. Choose the MX2 from the *Name – Model Name (Serial Number)* dropdown.  
   *Note*: Select the Refresh List button if the MX2 does not appear in the dropdown list.  
     
   
2. Select the *Upload MX2 Firmware* page from the list view on the left side of the diagnostic page.  
     
    
3. **Please complete steps a. – d. below twice.**
   1. Select the **Browse…** button.  
        
      
   2. Select the *MX2 1.0.11* file obtained during the .zip extract step above. Once uploaded, the following message will appear: *Validated. Upgrade will be applied during reboot.*



* 1. Select the **reboot** link within the message to navigate to the *Reboot* diagnostics page.



* 1. Click the **Reboot** button.  
       
     