TRX Firmware Install Instructions


Prior to installing, it is recommended to review the TRX Firmware Release Notes. The following are the instruction to install the firmware update.

# Accessing TRX Firmware


The firmware file is typically provided on a CD from DSI. If the firmware was e-mailed by DSI, save the attached firmware file to a location on your hard drive. Make a note of this location for **Step 8** below.

# Installing TRX Firmware


1. Launch **Ponemah**.
2. Select the **Hardware | PhysioTel Digital Telemetry Diagnostics…***Note*: If using *Ponemah v5.20*, select **Hardware | Diagnostics…**




1. Choose the CLC from the *Name – Model Name (Serial Number)* dropdown.  
   *Note*: Select the Refresh List button if the CLC does not appear in the dropdown list.




1. Select the *TRX Commands* link from the list view on the left side of the diagnostic page.




1. Select the TRX number for the firmware update from the **TRX No.***Note:*The TRX number corresponds to the CLC jack that the TRX is connected to.




1. Select **Upload Firmware** from the *TRX Command* dropdown.



1. Click the **Browse** button and navigate to the firmware file that was provided.



1. Navigate to the file location for the TRX firmware, select it, and click OK.
2. Click **Send**.



1. After the firmware has been updated, verify the firmware revision by clicking on the **TRX Status** link from the left menu.