# TRX Firmware Install Instructions

Prior to installing, it is recommended to review the TRX Firmware Release Notes. The following are the instruction to install the firmware update.

# Accessing TRX Firmware

The firmware file is typically provided on a CD from DSI. If the firmware was e-mailed by DSI, save the attached firmware file to a location on your hard drive. Make a note of this location for Step 8 below.

# Installing TRX Firmware

- Launch Ponemah.

- Select the Hardware | PhysioTel Digital Telemetry Diagnostics…
Note: If using Ponemah v5.20, select Hardware | Diagnostics…

- Choose the CLC from the Name – Model Name (Serial Number) dropdown.
Note: Select the Refresh List button if the CLC does not appear in the dropdown list.

- Select the TRX Commands link from the list view on the left side of the diagnostic page.

- Select the TRX number for the firmware update from the TRX No.
Note: The TRX number corresponds to the CLC jack that the TRX is connected to.

- Select Upload Firmware from the TRX Command dropdown.

- Click the Browse button and navigate to the firmware file that was provided.

- Navigate to the file location for the TRX firmware, select it, and click OK.

- Click Send.

- After the firmware has been updated, verify the firmware revision by clicking on the TRX Status link from the left menu.